DAILY BALANCE VERIFICATION SHEET                              SOP: 4FAC-3                                                                                     Version: 04

| Special Instructions |  | Verification Limits (± 0.5 %) |  |  |
| --- | --- | --- | --- | --- |
| 1. Ensure that the balance is stabilized at least for 30 minutes before use. |  |  | Lower Limit | Upper Limit |
| 2. Check that balance is levelled. |  | 1st weight 500 g | 499.5 g | 500.5 g |
| 3. Check that the balance reads zero prior to use. |  |  |  |  |
| 4. Place the calibrated mass pieces on the balance. |  | 2nd weight 1 kg | 999.95 g | 1.05 kg |
| 5. Record the date, actual readings per calibrated mass pieces and sign the verification sheet. |  |  |  |  |
|  |  | 3rd weight 10 kg | 9.95 kg | 10.05 kg |
| 6. If the readings are out of the following limits, the balance must NOT be used and reported to the QA Manager immediately. |  |  |  |  |

| Verification Data for ______________________ (month) ___________ (year) |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Date Verified | Stabilization |  | Level (YES / NO) | Actual verification readings |  |  | Readings within limit (YES / NO) | Cleaned | Verified by | Checked by |
|  | Time on | Time started |  | 1st weight | 2nd weight | 3rd weight |  |  |  |  |

ILCO Farming                                                                                                   Confidential     Page_____of_____